# Spectrometer Logger

Author: Yen-Chun Huang

Date: 06/24/2024

# Goal

The spectrometer logger is a python-based logging script to record and monitor the devices status all day.

# Python Requirements

OS: Windows

Minimum Version: Python 3.11

# Installation

Download logger program from the link. In the terminal, change the directory to the folder where the pyB12logger is. Then type in ‘pip install -e pyB12LOG’.

# General Usage of pyB12logger

The pyB12logger is used to log data/status from instruments. When first start the logger, some configuration files and log folders will be created automatically.

To start the logger, type in the ‘pyB12logger’ terminal window. You are allowed to start only one pyB12logger.

To stop the logger, type in the ‘pyB12logger stop’ terminal window.

There are some arguments that can be passed through pyB12logger. The arguments are stated in the below table. Please be advised that the options of startup and desktop are only available in the administration mode.

| Arguments/Options | Description |
| --- | --- |
| {start, stop} | To start or stop logger. If no arguments or options present, it will be set to ‘start’ |
| -h --help | Show help message |
| -desktop {True, False} | Create desktop icons |
| -startup {True, False} | Enable or disable pyB12logger running when startup |
| -debug {True, False} | Start pyB12logger with console window |

# Log Data and Debug Log

The device logging files are in C://Users/Public/B12TLOG.The format of file name is ‘log_YYYYMMDDHHMM.csv’. The file size is less than 100 kB by default. The debug_log.txt is saved in this folder. The file size can be configured by a user.

# Logger Configuration

The default configuration file is in C://Users/Public/B12TLOG_Config/config.cfg. A user can modify or add configuration to match the following format.

[name_of_device]

configuration_key = value

…

variable1 = command1, alias = var1, ….

Variable2 = command2, alias = var2, ….

The first part of the format is for device connection. The device configuration keys and value are shown below. Only ‘device_status’ and ‘address’ are necessary to be provided. If a configuration key is not provided, a default value will be set for the key. DO NOT use quote mark for any of the following values. The logger handles the value type internally.

| Configuration Key | Default Value | Comments |
| --- | --- | --- |
| device_status | False | If True, the logger will try to build connection with this device. |
| address | None | The pyVisa address of a device |
| id_command | *IDN? | The command to check connection (String) |
| baud_rate | 9600 | The baud rate for connection |
| termination | \n | The terminations for both read and write |
| data_bits | 8 | The data bit for connection |
| flow_control | None | The flow control for connection |
| parity | None | The parity for connection |
| stop_bits | None | The stop bits for connection |
| delimiter | , | The delimiter to separate string to a list to get data value |
| index | 0 | The index of data value in a list converted from a string by the set delimiter |

The second part is for variables and commands. The logger first saves variables to each log as a header (1st row). Secondly, it sends commands to device, queries data and processes it based on the configuration, and then saves the data to each log (after 1st row). The format of log will be:

Date, Time, variable1, variable2, ….

2024-06-24, 12.29.05, 067, 033, ….

2024-06-24, 12.29.15, 067, 033, ….

….

In each variable-command line, you can also add some optional arguments. Please be notified that any arguments that you add have effects on pyB12monitor only and won’t affect the behavior of the pyB12logger. These arguments are separated using ‘, ’ (a comma and a space). The available arguments are shown below.

| Arguments | Functions | Example |
| --- | --- | --- |
| alias | Change the displayed name in monitor. Space is allowed. | alias = He Pressure |
| min | If a read value is lower than min, the monitor will show a warning message | min = 10 |
| max | If a read value is higher than max, the monitor will show a warning message | max = 25 |
| static | If a read value is equal to static, the monitor will show a warning message | static = 17 |
| bits | If it is specified and not equal to 0, indicators will be added to the monitor with the number of bits. It will also disable the display of the variable and convert read value to 16-bits data. | bits = 9 |
| bit_static | if bits is specify and a read value is not equal to bit_static, the monitor will show a warning message | bit_static = 301 |
| indicators | The list of indicator names. Please be aware that names must have single quote marks and separated by a comma (no space), and the length of list must be greater than or equal to specified bits. If there is a star ‘*’ before a name, it means the read Boolean value will be negated (Logical Not). | indicators = [‘status1’,’*status2’,’*status3’,….] |

# pyB12monitor

The pyB12monitor is used to plot historical data in real-time or static way. To start the monitor, type one of the following commands to start.

| Commands | Description |
| --- | --- |
| pyB12monitor | Start the plotting |
| pyB12monitor_debug | Start the plotting with console |

# Usage of pyB12monitor

Warnings: showing the warnings of devices based on the given ‘min’, ‘max’ or ‘static’ in configuration file.

Clears: clear the warning messages.

Window Length: the number of data points shown in the figure.

Hidden: Hide the selected items in shown list.

Shown: Show the selected items in hidden list

Status: showing the status of devices and logger. Green is ok and Red is warning.

Start Time & End Time: Both can be optional. Click ‘OK’ to review the static data within the given time range and click ‘Reset’ to review the live figure.

File (menu bar): select file to plot. Click ‘Reset’ to review the live figure.

Please be notified that the settings will be automatically saved and retrieved when pyB12monitor restarts.